## Slide 1: QUS: Scatterer Size
BME 844 – HW 7
Will Long

## Slide 2: Scatterer Size and Backscatter Properties
Faran Model [1]
  For a >> λ, interface with f0, i.e. frequency independent
  For a << λ, Rayleigh scatterers with frequency dependence of f4
Size of scatterers or ESD between these extremes can be inferred from spectrum of backscattered echoes
[1] Faran Jr, James J. "Sound scattering by solid cylinders and spheres." The Journal of the acoustical society of America 23.4 (1951): 405-418.

## Slide 3: Measurement of Backscatter [2,3]
Normalized power spectral density of windowed echo signal (over -6 or -12 dB bandwidth)
Normalization of transducer/imaging parameters
  Area of transducer aperture
  Distance to gated volume
  Axial length of windowed RF segment
Normalization by f4 term (i.e. normalized by small scatterer condition a << λ)
Calibration stage to normalize transducer frequency dependence
  Deviation in frequency dependence is thus attributed to the scatterer size
[2] Insana, Michael F., and Timothy J. Hall. "Parametric ultrasound imaging from backscatter coefficient measurements: Image formation and interpretation." Ultrasonic Imaging 12.4 (1990): 245-267.
[3] Lizzi, F. L., et al. "Relationship of ultrasonic spectral parameters to features of tissue microstructure." Ultrasonics, Ferroelectrics, and Frequency Control, IEEE Transactions on 34.3 (1987): 319-329.
Notes: Calibration stage consists of a flat interface and serves to eliminate normalize out the frequency dependence of tdr and electronic components

## Slide 4: Diameter Estimation
Measured backscatter  Match model to spectrum  Extract corresponding scatterer size (least squares)
Form factor [2,4]
  Faran theory, spherical shell, FFSM, Gaussian models
  Analytical description of backscatter power spectrum based on a given scatterer size
[4]
Linear regression [3]
  Slope and intercept of best-fit line of power spectrum
LUT often used to final generate image [2]
[2] Insana, M. F., et al. "Parametric ultrasound imaging from backscatter coefficient measurements.”
[3] Lizzi, F. L., et al. "Relationship of ultrasonic spectral parameters to features of tissue microstructure.”
[4] Oelze, M. L., et al. "Characterization of tissue microstructure using ultrasonic backscatter: Theory and technique for optimization using a Gaussian form factor." The Journal of the Acoustical Society of America 112.3 (2002): 1202-1211.
Notes: Form factor models derived to relate scatterer size to the observed power spectrum
  Form factor is proportional to FT of autocorrelation function i.e. power spectrum for WSS process
Gaussian describes spheres whose acoustic impedance vary continuously

## Slide 5: Drawbacks [2,3,5]
Common assumptions:
  Discrete scatterers within uniform background
  Tissue homogeneity
  Backscatter echoes generated from a stationary process
  Behavior with frequency solely dependent on size
  Scatterers within ROI are of single size
Size estimates are dependent on frequency band
  Bias in estimates with different center frequency
  Larger bandwidth makes fit more robust to noise
Backscattered signal is attenuation dependent
  Local variations affect scatterer size prediction
Highly dependent on model selection
[2] Insana, M. F., et al. "Parametric ultrasound imaging from backscatter coefficient measurements.”
[3] Lizzi, F. L., et al. "Relationship of ultrasonic spectral parameters to features of tissue microstructure.”
[5] Nordberg, Eric P., and Timothy J. Hall. "Effective Scatterer Diameter Estimates for Broad Scatterer Size Distributions." Ultrasonic imaging (2014)
Notes:   Single scatterer size may not offer enough insight into structure
  Accuracy diminishes when scatterer size is more distributed
  Frequency changes may also be attributed to scatterer shape, elasticity, orientation
  Different tissues are modeled better by specific form factors

## Slide 6: Evaluation of Form Factor Models [5]
Accuracy of different single ESD form factors for modeling narrow and broad scatterer size distributions
  Gaussian and Faran models accurately describe narrow size distributions
  Exponential model provides better fit to broad scatterer size distributions
  Form factor description should not serve as a 1:1 indication of scatterer size, but rather an indirect descriptor of general scatterer source
[5] Nordberg, Eric P., and Timothy J. Hall. "Effective Scatterer Diameter Estimates for Broad Scatterer Size Distributions." Ultrasonic imaging (2014)
Notes:   How well different form factors fit provide insight into scatterer distribution in ROI
Difficulty transfering
ESD alone is not enough to characterize tumor malignancy

## Slide 7: Assessment of Tumor Malignancy
Thyroid carcinoma staging in mice [6]
  Smaller ESD in cancerous cells
  Efficacy in human model remains unknown
Breast cancer tumor staging [7]
  In conjunction with EAC and SAS, provides high sensitivity and specificity in tumor identification
  Tissue attenuation and patient to patient variation resulted in poor separability of ESD
Evaluation of chemotherapy [8]
  Differences in changes of EAC and ESD were statistically significant between responsive and non-responsive patients
[6] Lavarello, Roberto J., et al. "Characterization of thyroid cancer in mouse models using high-frequency quantitative ultrasound techniques." Ultrasound in medicine & biology 39.12 (2013): 2333-2341.
[7] Tadayyon, Hadi, et al. "Quantitative ultrasound characterization of locally advanced breast cancer by estimation of its scatterer properties." Medical physics 41.1 (2014): 012903.
[8] Sannachi, Lakshmanan, et al. "Non-invasive evaluation of breast cancer response to chemotherapy using quantitative ultrasonic backscatter parameters." Medical image analysis 20.1 (2015): 224-236.
Notes: Thyroid: not a clinical freq range, heterogenous attenuation profiles, aberration, blood flow, clutter
EAC = effective acoustic concentration
SAS = spacing among scatterers

## Slide 8: HIFU Monitoring [9]
Non-invasive thermal ablation via ultrasound
QUS capable of monitoring temperature elevations during ablation
  Source of contrast, however, is unclear
ESD variation with treatment was inconsistent between trials
Different exposure conditions resulted in different ESD behavior
Spatial variation lesion attenuation with HIFU are difficult to account for and will bias ESD estimates (<10 % bias)
[9] Ghoshal, Goutam, et al. "Quantitative ultrasound imaging for monitoring in situ high-intensity focused ultrasound exposure." Ultrasonic imaging 36.4 (2014): 239-255.

## Slide 9: Conclusions
Frequency dependence of backscatterer with varying scatterer sizes allows characterization of tissue types
Modeling backscatter spectrum via Faran theory or other form factors gives an indication of average scatterer size
Current work on tumor characterization, performance of form factor models, HIFU monitoring
Limitations in vivo
Notes: Limitations due to model selection, attenuation dependence, noise due to bandwidth and bias center frequency